Table 1a. Characteristics of Alarm Events

| Characteristic | All Alarms | Technical Alarms | Clinical Alarms |
| --- | --- | --- | --- |
| Events, n (%) | 193,852 (100%) | 134,739 (69.5%) | 59,113 (30.5%) |
| Alarm density (/bed/day) | 82.9 | 57.7 | 25.3 |
| Priority |  |  |  |
| ADVISORY, n (%) | 135,783 (70%) | 134,134 (99.6%) | 1,649 (2.8%) |
| WARNING, n (%) | 57,847 (29.8%) | 605 (0.4%) | 57,242 (96.8%) |
| CRISIS, n (%) | 222 (0.1%) | 0 (0%) | 222 (0.4%) |
| Dismissed alarms |  |  |  |
| Dismissed, n (%) | 5,759 (3%) | 914 (0.7%) | 4,845 (8.2%) |
| Time-to-dismiss, median [IQR], sec | 20 [10-49] | 206 [17-3025] | 18 [10-36] |
| Alarm duration |  |  |  |
| Duration, median [IQR], sec | 6 [1-20] | 4 [1-17] | 12 [4-23] |

The study was conducted in a surgical ward comprising 41 monitored beds (12 high-acuity, 29 general) over a 57-day observation period (July 15 – September 9, 2025). Technical alarms indicate device or signal artifacts (e.g., SpO2 probe disconnection, electrode failure, signal loss). Clinical alarms indicate physiological parameter deviations or arrhythmias. Alarm density was calculated as total events divided by the number of monitored beds and observation days. Time-to-dismiss was calculated only among dismissed alarms (silenced = TRUE). IQR, interquartile range.

Table 1b. Response Characteristics of Clinical Alarms Stratified by Priority Level

| Characteristic | ADVISORY | WARNING | CRISIS |
| --- | --- | --- | --- |
| Events, n (% of clinical alarms) | 1,649 (2.8%) | 57,242 (96.8%) | 222 (0.4%) |
| Events, n (% of ALL alarms) | 1,649 (0.9%) | 57,242 (29.5%) | 222 (0.1%) |
| Dismissed, n (%) | 61 (3.7%) | 4,626 (8.1%) | 158 (71.2%) |
| Time-to-dismiss, median [IQR], sec | 19 [12-69] | 19 [10-37] | 8 [6-12] |
| Duration, median [IQR], sec | 18 [10-19] | 11 [4-23] | 31 [9-37] |
| Ward zone |  |  |  |
| High-acuity zone, n (%) | 770 (46.7%) | 8,150 (14.2%) | 105 (47.3%) |
| General zone, n (%) | 879 (53.3%) | 49,092 (85.8%) | 117 (52.7%) |

Priority levels are defined by the bedside monitoring system: ADVISORY indicates a low-urgency physiological deviation; WARNING indicates a potentially serious condition requiring attention; CRISIS indicates a life-threatening condition requiring immediate intervention. Time-to-dismiss was calculated only among dismissed alarms. The proportion of each priority level relative to all alarms (including technical) reflects the signal-to-noise ratio faced by nursing staff. Ward zone classification: high-acuity zone comprised 12 beds (rooms 1101-A–E, Re-A/B, W10-01–05); general zone comprised 29 beds. IQR, interquartile range.

Table 2. Observed alarm silencing rates by priority level and alarm burden quartile

| Priority | Q1 (0–9) | Q2 (9–18) | Q3 (18–31) | Q4 (31–97) |
| --- | --- | --- | --- | --- |
| ADVISORY | 3.3% [2.3–4.9] (n= 777) | 2.8% [1.6–4.9] (n= 394) | 5.2% [3.1–8.5] (n= 271) | 4.8% [2.6–8.7] (n= 207) |
| WARNING | 10.6% [10.1–11.2] (n=13,921) | 8.4% [8.0–8.9] (n=14,329) | 7.6% [7.2–8.1] (n=14,455) | 5.8% [5.4–6.2] (n=14,537) |
| CRISIS | 70.4% [59.7–79.2] (n= 81) | 69.1% [56.0–79.7] (n= 55) | 76.9% [63.9–86.3] (n= 52) | 67.6% [50.8–80.9] (n= 34) |

Values are observed silencing rates (%) with 95% Wilson confidence intervals. Alarm burden quartiles represent the number of clinical alarms in the same ward in the preceding 30 minutes (Q1: 0–9, Q2: 9–18, Q3:18–31, Q4: 31–97 alarms/30 min). N per cell indicates the number of clinical alarms in that stratum. Total N = 59,113. ADVISORY alarms are underrepresented because most are generated by a small number of beds with vpersistent low-priority signals. CRISIS alarms maintained silencing rates of 68–77% across all burden quartiles, whereas WARNING silencing rates declined monotonically from 10.6% (Q1) to 5.8% (Q4).

Figure 1. Alarm silencing rate by alarm burden and priority level: evidence for resilient priority discrimination

Observed alarm silencing rates (%) with 95% Wilson confidence intervals, stratified by alarm priority (ADVISORY, WARNING, CRISIS) and alarm burden quartile (Q1–Q4). Alarm burden represents the number of clinical alarms in the same ward in the preceding 30 minutes (Q1: 0–9, Q2: 9–18, Q3: 18–31, Q4: 31–97 alarms/30 min). Lines connect median values within each priority group. N = 59,113 clinical alarms.

Table 3. Mixed-effects logistic regression: association between alarm burden, priority, and nurse silencing response

| Variable | OR | 95% CI | p-value |
| --- | --- | --- | --- |
| Alarm priority (ref: ADVISORY) |  |  |  |
| WARNING | 4.94 | 3.357 – 7.271 | <0.001 |
| CRISIS | 74.842 | 40.409 – 138.617 | <0.001 |
| Alarm burden × priority interaction |  |  |  |
| Per alarm increase (ADVISORY) | 1.011 | 0.993 – 1.029 | 0.236 |
| Per alarm increase × WARNING | 0.978 | 0.960 – 0.996 | 0.015 |
| Per alarm increase × CRISIS | 1 | 0.972 – 1.029 | 0.999 |
| Covariates |  |  |  |
| Zone: high-acuity (ref: general) | 1.461 | 0.971 – 2.198 | 0.069 |
| Shift: evening (ref: day) | 0.78 | 0.706 – 0.861 | <0.001 |
| Shift: night (ref: day) | 1.459 | 1.365 – 1.560 | <0.001 |
| Weekend (ref: weekday) | 1.141 | 1.065 – 1.222 | <0.001 |
| Ward: West 11 (ref: West 10) | 1.489 | 1.010 – 2.195 | 0.044 |
| Random effect |  |  |  |
| Bed-level variance | 0.302 |  |  |

Mixed-effects logistic regression model with bed-level random intercept. The outcome is nurse-initiated alarm silencing (yes/no). Alarm burden is defined as the number of clinical alarms in the same ward in the preceding 30 minutes, excluding the index alarm. Interaction terms between alarm burden and priority level test whether the effect of burden on silencing probability differs across priority strata. OR = odds ratio; CI = confidence interval. N = 59,113 clinical alarms (July 15 – September 9, 2025). AIC = 31,653.3.